## Slide 1: High-Field Instrumentation for EPR and DNP Spectroscopy
Thorsten Maly
Bridge12 Technologies, Inc.
ECT* Tensor Spin Observables, Trento, Italy, 07/10 – 07/14/2023

## Slide 2
Bridge12 – Modular, agile high/low-frequency microwave instruments and systems for academic research
Two large labs for EPR/NMR/DNP
In-house workshop

## Slide 3: Progress of DNP-NMR Spectroscopy in the Bio Community
BDPA/PS, 1993
bTbK, 2009
e > 250
e ~ 10
e > 180
TOTAPO, 2006

## Slide 4: Optimizing Polarizing Agents for 1H DNP-NMR Spectroscopy 9.4 T, 400 MHz 1H-NMR
Kubicki et al. Chimia , 2017
Notes: 1H polarization close to 50 % at 90 K

## Slide 5: High-Field EPR Spectroscopy
FAD Cofactor Anion in Glucose Oxidase
Weber et al., J. Phys. Chem. B, 2008
Pros
  g-Tensor resolution
  Increased sensitivity
Cons
  Demanding instrumentation
  Often not turn-key operation
MagLab
Notes: Example is an organic radical in a protein
Increased resolution good to understand electronic structure, study vicinity and dynamics (libration)
Predominantly home-built instrumentation

## Slide 6: State-of-the-Art – High-Field DNP/NMR Spectroscopy
DNP-NMR (MAS) Spectroscopy
  Temperature regime: 90 K and up
  Uses magic angle spinning (MAS), some static experiments
  Turn-key method/instrumentation
    Stable sources (gyrotron)
    Optimized microwave coupling to sample
  Most commonly used at 9.4 T (400 MHz 1H)
  Try-and-error approach to optimizing polarizing agents. Need more EPR data (T1e, T2e, T1n, J, …) at different magnetic field strengths
  Already good understanding of mechanisms, but more work required especially for DNP-NMR experiments under MAS conditions

## Slide 7: Irradiation Only with Isolator
Basic setup
Source protection using isolator
  Important when using solid-state sources
Isolator

## Slide 8: Irradiation Only with Circulator
Basic setup
Source protection using circulator
Reflected power is directed to load
Minimize standing waves
Circulator
Absorber
Circulator
3 or 4 port device
Directs microwave power
  Port 1 → Port 2
  Port 2 → Port 3
  Port 3 → Port 1

## Slide 9: Reflection Bridge Direct Detection
Basic setup for EPR detection
Source protection through circulator
Reflected power is directed to detector (need resonator)
No phase information (superposition of absorption and dispersion signal)
Thurber et al. “Low-Temperature Dynamic Nuclear Polarization at 9.4T with a 30mW Microwave Source.” Journal of Magnetic Resonance 2010
TEMPO @ 263 GHz
LNA
Detector
Notes: Could be a very simple EPR spectrometer, but you need a very well tuned microwave cavity (resonator)

## Slide 10: Superheterodyne Detection Phase Sensitive and Highest Overall Sensitivity
Most sophisticated setup for EPR detection
Maximum sensitivity (amplification at IF)
Create/Shape pulses at LO frequency
AMC
Offset
50 MHz
LO
IF
Add Pulse Forming Unit
Create/shape pulses at
LO frequency (e.g. x-band)

## Slide 11: Components
Transmission Lines
Excitation
  Microwave Sources
  Pulse Shaping
Detection
  cw or time domain
Discreet Components
  Circulators
  Isolators
  Power Splitters/Combiners

## Slide 12: Transmission Lines

## Slide 13: Transmission Lines
Coaxial
Cable
Waveguide
Corrugated
Waveguide
RF to 18-28 GHz depending on connector (N-type, SMA, 2.92 …)
< 28 GHz: low losses
> 95 GHz: high losses
(22.7 dB/m for WR-4.3)
Excellent performance
Expensive
Extreme low losses at high frequencies.
Can be direct machined or electroformed
Corrugated Waveguides

## Slide 14: 3D Printed Hollow Fibers Something to keep an eye on …
Hong et al., “Low-Loss Asymptotically Single-Mode THz Bragg Fiber Fabricated by Digital Light Processing Rapid Prototyping.” IEEE Transactions on Terahertz Science and Technology 8, no. 1 (January 2018): 90–99
~3dB/m

## Slide 15: Microwave Sources

## Slide 16: Microwave Sources - Tube Map
Vacuum Electronic Devices (VED)
  Gyrotron
  Klystron (EIO, EIK ..)
  Traveling Wave Tube (TWT)
  Backward Wave Oscillator (BWO)
Gyrotron is the only demonstrated high-power source at frequencies > 263 GHz
Watch Monica Blank’s talk: “Millimeter Sources for NMR/DNP” on YouTube

## Slide 17: Gyrotron-Based DNP-NMR Spectroscopy
Rosay, M., M. Blank, and F. Engelke. “Instrumentation for Solid-State Dynamic Nuclear Polarization with Magic Angle Spinning NMR.” J Magn Reson 264 (March 2016): 88–98

## Slide 18: Solid-State Microwave Sources
Principle of Operation
  Start at X-Band (8 – 12 GHz)
  Use series of amplification and multiplication steps (AMC)
  Key player: Virginia Diodes
Benefits
  Truly frequency agile
  Cost-effective
  Small footprint
  Turn-key
  No additional requirements (cooling water, 3-phase power, air …)
Limitations
  Low power (but … steadily increasing) up to 800 mW @ 140 GHz > 250 mW @ 263 GHz > 100 mW @ 380 GHz
Amplification Multiplication Chains (AMCs)
(ww.vadiodes.com)
Notes: Power depends on the bandwidth

## Slide 19: Source Comparison After Monica Blank’s talk on the MIT zoominar
 | Solid-State | VED (EIO/EIA) | Gyrotron
Size/Weight | Compact | Compact | Large
Facility Requirements | Minimal | Cooling water, HV | Cooling water, air HV
Output power (for DNP) | < 250 mW | 1 – 10 W | > 50 W
Cost | $ | $$ | $$$
Cost/Watt | $$$ | $$ | $

## Slide 20: Pulse Shaping
PELDOR/DEER Spectroscopy
  Minimize area of overlap to
    Minimize detection artifacts
    Acquire more accurate/reliable distance distributions
Endeward et al., JMR, 2023
Pulse Shaping
  Tailor the pulse excitation profile for optimum performance

## Slide 21: Pulse Shaping
PFU
RCVR
PRELIMINARY RESULTS

## Slide 22: Discreet Microwave Components

## Slide 23: High-Frequency Circulator
Low Frequencies (< 100 GHz)
Waveguide (or coaxial) transmission
Low insertion losses
Many discrete components
  Circulators
  Attenuators
  Phase shifters
94 – 96 GHz
1.5 dB loss
Use low-low Quasi-Optics for Beam Manipulation

## Slide 24: What are Quasi-Optics (QO)?
Free-space Gaussian Beam (TEM00)
Manipulate beam using combination of mirrors, grids, and l/4-plates (or equivalent)
Extremely low losses
Every discreet microwave component has a QO equivalent
95 GHz launcher, 97.8 % HE11 –TEM00 coupling
Each QO component has a quadratic footprint, arranged on a grid
Distance between focusing elements is constant
Polarization controlling elements are placed in between pair of mirrors
  Gaussian Telescope
Notes: Magnification: Unity
Frequency Independent

## Slide 25: Polarization Transforming Reflector (PTR)
Introduce phase difference to change polarization
Convert linear to circular polarization or any in between
Low insertion loss, high isolation
Compact design
Insertion Loss: 0.80 dB
Isolation: 30.9 dB

## Slide 26: 140/350 GHz Dual-Frequency EPR Spectrometer
PTR-based QO Duplexer
140/350 GHz EPR Spectroscopy
  Max. magnetic field: 14.5 T
  Temperature regime: 1.95 to 300 K
  Liquid cryogen free operation
  Microwave power
    140 GHz: > 500 mW
    350 GHz: > 100 mW
  Operation modes:
    CW
      lock-in detection
    Time-domain
      Discreet and AWG generated pulses
  Completely automated operation

## Slide 27: 140/350 GHz Dual-Frequency EPR Spectrometer
CW
Pulsed
PRELIMINARY RESULTS

## Slide 28
THANK YOU
DNP Literature Blog: blog.bridge12.com
Twitter: @thmaly
Email: tmaly@bridge12.com

## Slide 29: Building an EPR Community Through OS Hard- and Software
pip install dnplab
import dnplab as dnp
data = dnp.load("topspin/1")
DNPLab is created by Bridge12, the Han Lab (UCSB) and the Franck Lab (Syracuse).
data = dnp.apodize(data,lw=100)
data = dnp.fourier_transform(data)
data = dnp.autophase(data)
dnp.fancy_plot(data)
dnp.show()
Python-based
Object-oriented
Open-source
Modulation
Coils
Resonator
Single PCB
Hardware < 50 $
Easily assembled
Teach EPR spectroscopy
Teach instrumentation
OpenPCBLGR
Find it on GitHub
First fully open-source X-Band Loop Gap Resonator

## Slide 30: Q-Band (35 GHz) EPR Spectrometer
Q-Band EPR Spectrometer
  Max. magnetic field: 3.5 T
  Temperature regime: 1.95 to 300 K
  Liquid cryogen free operation
  Microwave power
    34 GHz: > 100 W
  Time Domain Operation
    Fully AWG controlled
  Completely automated operation

## Slide 31: QO Duplexer using a PTR
Excitation
PTR changes beam polarization from linear to circular (e.g. RH)
Detection
Reflected beam has opposite polarization (e.g LH)
PTR changes polarization back to linear (but orthogonal to incident beam)
140 GHz
263 GHz

## Slide 32: Compact Quasi-Optical Bridges for Beam Manipulation Application: Gyrotron-based DNP-NMR Spectroscopy
QO bridge for gyrotron-based DNP-NMR spectrometers
Continuous, remote-controlled attenuator
Universal polarizer
Shutter to create ms pulses
Chen et al., Rev. Sci. Inst, 2022 https://doi.org/10.1063/5.0036292
Polarizing Transforming Mirror and Universal Polarizers

## Slide 33: Cube Layout of Quasi-Optics Systems
Each QO component has a quadratic footprint
All components are arranged on grid
Mirror, grid holders are designed as half cube elements
Distance between focusing elements is constant
Polarization controlling elements are placed in between pair of mirrors
  Gaussian Telescope
Lesurf, J. C. G. Millimetre-Wave Optics, Devices, and Systems. Bristol, England ; New York: A. Hilger, 1990.
Magnification: Unity
Frequency Independent
Gaussian Telescope

## Slide 34: What is a Gaussian Beam?
Gaussian beam mode: TEM00
Efficient coupling of the TEM00 mode to the HE11 mode
Description is extensively used in optical spectroscopy
Fully characterized by its beam waist radius (w) and radius of curvature (R)
At the origin w = w0 and R = (flat beam front)
Goldsmith, Paul F. Quasioptical Systems: Gaussian Beam Quasioptical Propagation and Applications. IEEE Press/ Champman & Hall Publishers, 1998.
Beam waist radius expands while propagating
Expansion depends on frequency
Each transition HE11 to TEM00 introduces ~2 % loss

## Slide 35: Components of a Quasi-Optical System
Launchers
  Input/Output Port
Mirrors
  Focusing Mirrors
  Flat Mirrors
(Lenses)
Wire Grids
  Polarization Control
Wave Plates
  Quarter Wave Plates
  Faraday Rotator
Beam Splitters
Absorbers

## Slide 37: QO Duplexer – Faraday Rotator For Superheterodyne Detection
Excitation
Source (A5) protected by grid (B3) from reflected power
Faraday rotator located in C3, rotates polarization by 45º
Mirror (D3) can be used to direct beam out of plane
Sample (D1) is illuminated with a linearly polarized beam
Induction Mode Detection
Power reflected by sample (D1) passes grid (C3)
Beam has elliptical polarization
Some power proportional to absorption is reflected by grid C3, re-focused (C5) and detected E5
Reflection Mode Detection
Power reflected by sample (D1) passes grid (C3)
Faraday rotator rotates polarization again by 45º
Beam is reflected by grid B1
Beam focused (C4) and directed to detector (D4)
Change in intensity corresponds to EPR resonance

## Slide 38: Off-Axis Mirrors
Flat Mirror
Focusing Mirror
Rooftop Mirror
Jones Matrix
Ray Propagation
Change propagation direction by 90º
(Re)-focus
expanding beam
Reverse beam
propagation direction

## Slide 39: Image Distortion in Mirror Pairs
Trans configuration yields more symmetric beam
Cis configuration has lower cross-polarization content

## Slide 40: DNP is a Proven Method to Enhance NMR Signal Intensities
Temperature-Jump
Polarize solid at high fields
Melt sample with laser flash
Record solution-state NMR
Freeze sample, repeat cycle
Dissolution
Polarize solid at low field
Melt sample with hot solvent
Transfer sample
Record solution-state NMR
Pump/Shuttle
Polarize solution at low fields
Shuttle sample to high field
Record solution-state NMR
Shuttle sample back, repeat cycle
Direct Polarization
Polarize sample directly
Solids: Straightforward approach
Solution: “Easy” at low fields, “Difficult” at high fields
Notes: DNP enhances NMR signal intensities. The method is successfully employed in solid-state NMR spectroscopy
We heard about all of these different methods.

## Slide 41: High-Field EPR/DNP Spectroscopy – Why Bother?
High-Field EPR Spectroscopy
  Paramagnetic bio-macromolecules (electron transfer proteins)
  Material science (catalysis, battery materials)
High-Field DNP-NMR Spectroscopy
  Sensitivity enhancement of NMR experiments
  Surface characterization
  Domain size characterization